A PROJECT REPORT ON

IoT Health Monitoring System (Heart Rate, SpO₂, Body Temperature)

IoT project using ESP32 with Bluetooth Low Energy and Wi-Fi (ThingSpeak)

Project: P08 · Advanced level

Peripherals: MAX30102 pulse oximeter, DS18B20 temperature probe, SSD1306 OLED, Active buzzer

Submitted by: ______________________________

Under the guidance of: ______________________________

Department / Institution: ______________________________

October 2026

# Certificate

This is to certify that the project report entitled “IoT Health Monitoring System (Heart Rate, SpO₂, Body Temperature)” is a bonafide record of the work carried out by ______________________ (Roll No. ________) in partial fulfilment of the requirements for the award of ______________________ in ______________________ during the academic year ________.

Project Guide: ______________________          Head of Department: ______________________

Date: ____________          Place: ____________

# Acknowledgement

I express my sincere gratitude to my project guide ______________________ for continuous guidance and support, to the Head of the Department ______________________, and to all faculty members and friends who helped me complete this project successfully.

# Contents

Abstract	6

1. Introduction	7

1.1 Problem Statement	7

1.2 Objectives	7

1.3 Scope of the Project	7

1.4 Organisation of the Report	7

2. Existing System and Proposed System	8

2.1 Existing Systems	8

2.2 Proposed System	8

2.3 Why Both BLE and Wi-Fi?	8

3. System Design	9

3.1 Block Diagram	9

3.2 System Architecture	9

3.3 Hardware Components and Cost	9

3.4 ESP32 DevKit V1	10

3.5 Working Principle of Components	10

4. Circuit Design	11

4.1 Circuit Diagram	11

4.2 Pin Connections	11

4.3 Interfacing and Power Notes	12

5. Data Flow Diagrams	13

5.1 DFD Level 0 (Context Diagram)	13

5.2 DFD Level 1	14

5.3 Data Dictionary	15

6. Software Design and Implementation	16

6.1 Development Environment	16

6.2 Firmware Architecture	16

6.3 Working of the System	16

6.4 Flowchart	17

6.5 BLE Communication Protocol	18

6.6 Cloud Integration (ThingSpeak)	18

6.7 Configurable Settings	19

7. Setup and Deployment Guide	20

7.1 Hardware Assembly	20

7.2 Flashing the Firmware	20

7.3 Cloud Setup (ThingSpeak)	21

7.4 BLE App Setup (Web Bluetooth Console)	21

7.5 First-Run Checklist	22

8. Testing and Results	23

8.1 Expected Output	23

8.2 Troubleshooting	23

9. Discussion	24

9.1 Advantages	24

9.2 Limitations	24

9.3 Safety Precautions	24

9.4 Security Considerations	24

9.5 Future Scope	24

10. Conclusion	25

References	26

Appendix A: Project Source Code	27

Appendix B: Shared Library Code	31

# Abstract

This project measures three basic vital signs — heart rate, blood-oxygen saturation (SpO₂) and body temperature — and makes them available locally and remotely. A MAX30102 optical sensor captures red and infrared photoplethysmography (PPG) signals from a fingertip; the ESP32 samples them without blocking at 25 Hz into a sliding four-second window and runs Maxim's reference algorithm every second to estimate heart rate and SpO₂. A waterproof DS18B20 probe measures temperature. Readings appear on an OLED, stream to a phone over BLE and are uploaded to ThingSpeak every 20 seconds while a finger is on the sensor. Configurable limits trigger a buzzer and a phone alert. The system is an educational prototype and not a medical device.

Keywords: ESP32, Internet of Things, Bluetooth Low Energy, Wi-Fi, ThingSpeak, MAX30102 pulse oximeter, DS18B20 temperature probe, SSD1306 OLED, Active buzzer.

# 1. Introduction

The Internet of Things (IoT) connects physical devices — sensors and actuators — to networks so that data can be collected, analysed and acted on automatically. The ESP32 from Espressif Systems has made IoT affordable: a single ₹350 board combines a dual-core 240 MHz processor, Wi-Fi and Bluetooth Low Energy. This project uses both radios at once: BLE for direct, local interaction with a phone and Wi-Fi for cloud connectivity through ThingSpeak, a free IoT/cloud platform.

## 1.1 Problem Statement

Elderly people, patients recovering at home and people in rural areas often need periodic monitoring of pulse, oxygen and temperature, and caregivers may be far away. Clinical monitors are expensive, and stand-alone finger oximeters do not share or log their readings.

## 1.2 Objectives

- Measure heart rate and SpO₂ from PPG signals with a MAX30102.

- Measure body temperature with a DS18B20 probe.

- Detect finger presence and produce stable, averaged readings.

- Alert locally and on a phone when values leave configurable limits.

- Log readings to ThingSpeak so caregivers can view trends remotely.

## 1.3 Scope of the Project

Spot checks and short monitoring sessions at rest, for education and awareness. Continuous wearable use, motion tolerance and clinical accuracy are out of scope.

## 1.4 Organisation of the Report

Chapter 2 reviews existing approaches and describes the proposed system. Chapter 3 covers system design and components, Chapter 4 the circuit, Chapter 5 the data flow diagrams, Chapter 6 the software, Chapter 7 the step-by-step setup of hardware, cloud and app, Chapter 8 testing, Chapter 9 discussion and Chapter 10 the conclusion. Complete source code is given in the appendices.

# 2. Existing System and Proposed System

## 2.1 Existing Systems

Many MAX30102 projects use blocking loops that freeze the microcontroller for four seconds per reading, which breaks Wi-Fi/BLE responsiveness; others compute heart rate only. Results are rarely logged with finger detection, so empty readings pollute the data.

## 2.2 Proposed System

The proposed system improves on these in the following ways:

- Non-blocking FIFO reading with a sliding window: a new estimate every second.

- Finger detection by IR level; uploads only valid readings.

- Four-sample moving average of heart rate.

- Configurable alert limits stored in flash and changed from the phone.

## 2.3 Why Both BLE and Wi-Fi?

Wi-Fi gives the device internet access for remote monitoring, history and control from anywhere, but it depends on a router and an internet connection, and the device needs the Wi-Fi password before it can connect. BLE needs no infrastructure: any modern phone can connect directly within about 10-30 m. Using both gives:

- Provisioning — Wi-Fi credentials and cloud keys are sent from the phone over BLE and stored in flash; nothing secret is hard-coded or uploaded to GitHub.

- Resilience — local monitoring and control keep working when the internet or router is down.

- Remote access — the cloud dashboard works from anywhere when the internet is available.

- Low cost — no extra hardware; both radios are built into the ESP32.

The ESP32 shares one 2.4 GHz radio between Wi-Fi and Bluetooth by time-division coexistence. For this to work the Wi-Fi modem-sleep mode must stay enabled, which the firmware ensures. NimBLE, a lightweight BLE stack, is used instead of the default Bluedroid stack to leave more RAM for TLS (HTTPS/MQTTS) connections.

# 3. System Design

## 3.1 Block Diagram

Figure 1: Block diagram of the Health Monitoring System

The ESP32 is the central controller. Input devices (blue) feed measurements to it, output devices (amber) are driven by it, and two communication links (green) connect it to the user's phone over BLE and to the ThingSpeak cloud over Wi-Fi. The board is powered from 5 V through its USB connector or VIN pin.

## 3.2 System Architecture

The system follows the standard three-layer IoT architecture:

- Perception layer — MAX30102 pulse oximeter, DS18B20 temperature probe, SSD1306 OLED, Active buzzer connected to the ESP32 GPIO, ADC, I2C, SPI or UART.

- Network layer — BLE 4.2 GATT server (Nordic UART Service) and Wi-Fi 802.11 b/g/n with TLS-encrypted HTTPS REST to ThingSpeak.

- Application layer — the Web Bluetooth console on the phone and the ThingSpeak dashboard / database.

## 3.3 Hardware Components and Cost

| # | Component | Qty | Specification | Cost |
| --- | --- | --- | --- | --- |
| 1 | ESP32 DevKit V1 (ESP-WROOM-32, 30-pin) | 1 | Dual-core 240 MHz, Wi-Fi + BLE, 4 MB flash | ₹350 |
| 2 | Breadboard 830-point + jumper wires (M-M, M-F) | 1 | Prototyping and connections | ₹150 |
| 3 | Micro-USB data cable | 1 | Power + programming | ₹60 |
| 4 | MAX30102 pulse oximeter & heart-rate module | 1 | Red + IR LEDs, 18-bit ADC, I2C 0x57 | ₹150 |
| 5 | DS18B20 waterproof temperature probe | 1 | −55 to +125 °C, ±0.5 °C, 1-Wire | ₹90 |
| 6 | 4.7 kΩ resistor | 1 | 1-Wire pull-up | ₹2 |
| 7 | 0.96" SSD1306 OLED 128x64 (I2C) | 1 | Address 0x3C | ₹170 |
| 8 | Active buzzer 5 V | 1 | Alert | ₹15 |
|  | Total (approx.) |  |  | ₹987 |

Table 1: Bill of materials with approximate Indian prices (Robu.in / Amazon.in / local market, 2026)

Prices are indicative and vary by seller. The breadboard, cable and ESP32 can be reused across all ten projects, so the incremental cost of this project is about ₹427.

## 3.4 ESP32 DevKit V1

| Feature | Value |
| --- | --- |
| Module | ESP-WROOM-32 (ESP32-D0WD), dual-core Xtensa LX6 up to 240 MHz |
| Memory | 520 KB SRAM, 4 MB SPI flash |
| Wireless | Wi-Fi 802.11 b/g/n 2.4 GHz; Bluetooth 4.2 BR/EDR + BLE |
| I/O | 25 usable GPIO on the 30-pin board, 12-bit ADC (ADC1: GPIO32-39), 2 × DAC, PWM on any output |
| Interfaces | 3 × UART, 2 × I2C, 3 × SPI, I2S, touch, RMT |
| Supply | 5 V via USB/VIN, on-board 3.3 V regulator (AMS1117) |
| Logic level | 3.3 V — inputs are NOT 5 V tolerant |

Table 2: Key specifications of the ESP32 DevKit V1

## 3.5 Working Principle of Components

### Photoplethysmography

Arterial blood volume in the fingertip changes with every heartbeat, which changes how much light is absorbed. The MAX30102 shines red (660 nm) and infrared (880 nm) light and measures the reflected light. The periodic (AC) part gives the pulse; the time between peaks gives heart rate.

### SpO₂ ratio

Oxygenated haemoglobin absorbs more IR and less red light than deoxygenated haemoglobin. The ratio R = (AC_red/DC_red)/(AC_IR/DC_IR) is mapped to SpO₂ with an empirical curve; Maxim's algorithm uses SpO₂ ≈ −45.06·R² + 30.354·R + 94.845 via a lookup table.

### DS18B20

A digital thermometer on Dallas 1-Wire: one data line with a 4.7 kΩ pull-up. A 12-bit conversion takes 750 ms; the firmware requests a conversion and reads the result on the next pass instead of waiting.

### OLED & buzzer

The OLED shows bpm, SpO₂, temperature and a blinking heartbeat dot; the buzzer chirps while an alert is active.

# 4. Circuit Design

## 4.1 Circuit Diagram

Figure 2: Circuit diagram of the Health Monitoring System

## 4.2 Pin Connections

| Module | Module pin | ESP32 connection | Notes |
| --- | --- | --- | --- |
| MAX30102 | VIN | 3.3 V (3V3 pin) |  |
| MAX30102 | SDA | GPIO21 |  |
| MAX30102 | SCL | GPIO22 |  |
| MAX30102 | GND | GND |  |
| SSD1306 OLED | VCC | 3.3 V (3V3 pin) |  |
| SSD1306 OLED | SDA | GPIO21 |  |
| SSD1306 OLED | SCL | GPIO22 |  |
| SSD1306 OLED | GND | GND |  |
| DS18B20 probe | Red VDD | 3.3 V (3V3 pin) |  |
| DS18B20 probe | Yellow DQ | GPIO4 | 4.7 kΩ pull-up between DQ (GPIO4) and 3.3 V |
| DS18B20 probe | Black GND | GND |  |
| Active buzzer | + | GPIO18 |  |
| Active buzzer | - | GND |  |

Table 3: Pin connection table

## 4.3 Interfacing and Power Notes

- MAX30102 and OLED share SDA/SCL (different addresses, 0x57 and 0x3C).

- Some purple MAX30102 boards need the 1.8 V/3.3 V solder jumper set to 3.3 V for I2C pull-ups.

- Only ADC1 pins (GPIO32-39) are used for analog inputs, because ADC2 cannot be read while Wi-Fi is active.

- Strapping pins GPIO0, 2, 12 and 15 are avoided for inputs so the board always boots normally.

- All grounds are common. 5 V signals are reduced to 3.3 V with resistor dividers before reaching the ESP32.

# 5. Data Flow Diagrams

Data flow diagrams (DFDs) show how data moves through the system without showing control logic or timing. External entities are drawn as rectangles, processes as numbered circles, data stores as open boxes and data flows as labelled arrows.

## 5.1 DFD Level 0 (Context Diagram)

Figure 3: DFD Level 0 — context diagram

At level 0 the whole Health Monitoring System is a single process. It receives PPG red/IR, temperature from the Patient (fingertip, skin), commands and configuration from the user's phone, and no commands from the cloud. It sends vitals display, alert chirp to the OLED + buzzer, live data to the phone and HR, SpO₂, temp, alert every 20 s to ThingSpeak.

## 5.2 DFD Level 1

Figure 4: DFD Level 1 — main processes and data stores

| Process / store | Description |
| --- | --- |
| 1.0 Acquire sensor data | Reads PPG red/IR, temperature from the Patient (fingertip, skin); filters/averages the raw values. |
| 2.0 Process & decide | Computes HR, SpO₂, alerts; applies thresholds and rules; generates events. |
| 3.0 Actuate & alert | Drives the OLED + buzzer (vitals display, alert chirp). |
| 4.0 BLE communication | Sends JSON frames and events to the phone; receives and authenticates commands. |
| 5.0 Cloud sync | Sends HR, SpO₂, temp, alert every 20 s to ThingSpeak. |
| D1 Settings (NVS) | Wi-Fi credentials, cloud keys, PIN, thresholds — survive reboots. |
| D2 PPG sample window (RAM) & limits (NVS) | Stores 100 red/IR samples; returns window, limits. |

Table 4: Description of level-1 processes and data stores

## 5.3 Data Dictionary

| Field | Meaning | Unit |
| --- | --- | --- |
| hr | Heart rate | bpm |
| spo2 | Oxygen saturation | % |
| temp | Body temperature | °C |
| finger | Finger detected | bool |
| alert | Alert active | bool |
| msg | Alert details | — |

Table 5: BLE data frame fields

| Cloud object | Content / usage |
| --- | --- |
| Field 1 | Heart rate (bpm) |
| Field 2 | SpO₂ (%) |
| Field 3 | Temperature (°C) |
| Field 4 | Alert flag (0/1) |
| Status | Alert text or 'normal' |

Table 6: ThingSpeak data objects

# 6. Software Design and Implementation

## 6.1 Development Environment

| Item | Details |
| --- | --- |
| IDE / build | Arduino IDE 2.x or arduino-cli; GitHub Actions builds firmware in the cloud |
| Board package | esp32 by Espressif Systems 3.x (tested with 3.3.12) |
| Board settings | ESP32 Dev Module, Partition Scheme: Minimal SPIFFS (1.9MB APP with OTA/190KB SPIFFS) |
| Common libraries | NimBLE-Arduino 2.x, ArduinoJson 7.x, PubSubClient 2.8 |
| Project libraries | SparkFun MAX3010x Pulse and Proximity Sensor Library, OneWire, DallasTemperature, Adafruit SSD1306, Adafruit GFX Library |
| Phone app | Web Bluetooth console (Chrome on Android / desktop, Bluefy on iPhone); nRF Connect as alternative |
| Cloud | ThingSpeak (free tier) |

Table 7: Software tools

The default 1.2 MB app partition is too small for BLE + Wi-Fi + TLS. Select Minimal SPIFFS (1.9MB APP) in Tools → Partition Scheme, otherwise the upload fails with 'Sketch too big'.

## 6.2 Firmware Architecture

The firmware is split into three layers so that the ten projects share tested code:

- IoTCore.h (shared) — Wi-Fi station with auto-reconnect, NimBLE GATT server with the Nordic UART Service, line-based command parser shared by BLE and USB serial, PIN authentication, settings in NVS (Preferences), NTP time in IST, JSON helpers.

- Cloud.h (shared) — ThingSpeak HTTPS client, Adafruit IO MQTT/TLS client with a 28 messages/minute rate limiter, Supabase PostgREST client (insert, update, select).

- P08_Health_Monitoring_System.ino — sensor drivers, algorithms, automation rules and the project's own commands.

The main loop is non-blocking: every task is scheduled with millis() timers, BLE writes are queued from the BLE callback into a FreeRTOS queue and executed in loop(), so the BLE stack never waits for slow operations such as HTTPS requests.

## 6.3 Working of the System

The MAX30102 is configured for 100 samples/s with 4-sample averaging (25 effective samples/s), 411 µs pulse width and an 18-bit range. In every loop pass the firmware drains the sensor FIFO. IR above 50,000 counts means a finger is present; removing the finger resets the buffers.

Samples enter a 100-sample (4 s) window. Once full, every 25 new samples (1 s) the window is passed to maxim_heart_rate_and_oxygen_saturation(). Valid heart rates (30-220 bpm) are averaged over the last four estimates; valid SpO₂ values (70-100 %) are shown directly. The DS18B20 is read once per second without blocking.

Alerts compare the values with hrlow/hrhigh, spo2low and temphigh. An alert sends a 'vital_alert' event and chirps the buzzer. Every 20 s, if a finger is present and heart rate is valid, HR, SpO₂, temperature and the alert flag are uploaded to ThingSpeak.

## 6.4 Flowchart

Figure 5: Firmware flowchart

## 6.5 BLE Communication Protocol

The device is a BLE peripheral named IOT-P08-XXXX (XXXX = last bytes of the MAC address). It exposes the Nordic UART Service, which most BLE apps understand:

| Item | UUID | Properties |
| --- | --- | --- |
| Service (NUS) | 6E400001-B5A3-F393-E0A9-E50E24DCCA9E | Primary |
| RX — phone → ESP32 | 6E400002-B5A3-F393-E0A9-E50E24DCCA9E | Write, Write without response |
| TX — ESP32 → phone | 6E400003-B5A3-F393-E0A9-E50E24DCCA9E | Notify, Read |

Table 8: GATT service and characteristics

Commands are plain text lines (command:argument). Replies are one JSON object per line ending in a newline; long lines are split into MTU-sized notifications and re-joined by the app. Commands that change anything require auth:<PIN> first on each connection (default PIN 1234, change it with set:pin=XXXX).

| Command | PIN needed | Action |
| --- | --- | --- |
| auth:<PIN> | — | Unlock for this connection |
| status | No | Status JSON: Wi-Fi, IP, RSSI, heap, uptime, time |
| wifi:<SSID>,<password> | Yes | Save Wi-Fi credentials and connect |
| set:<key>=<value> | Yes | Save a setting (cloud keys, thresholds) |
| get:<key> | Yes | Read a setting (secrets masked) |
| reboot / forget | Yes | Restart / erase all settings |
| status | No | Device status (the project has no actuator commands) |

Table 9: BLE commands (common commands first, then project commands)

Example data frame sent by the device:

{"hr":76,"spo2":98,"temp":36.6,"finger":true,"alert":false,"type":"data"}

## 6.6 Cloud Integration (ThingSpeak)

Data is sent with an HTTPS GET request: https://api.thingspeak.com/update?api_key=KEY&field1=…&status=…. ThingSpeak returns the new entry number, or 0 if the update was rejected (for example when sent within 15 s of the previous one). The free licence allows 3 million messages per year and one update every 15 s per channel, so a 20 s interval leaves a safety margin.

| Object | Content |
| --- | --- |
| Field 1 | Heart rate (bpm) |
| Field 2 | SpO₂ (%) |
| Field 3 | Temperature (°C) |
| Field 4 | Alert flag (0/1) |
| Status | Alert text or 'normal' |

Table 10: ThingSpeak configuration

## 6.7 Configurable Settings

| Key | Default | Meaning |
| --- | --- | --- |
| tskey | — | ThingSpeak Write API Key |
| hrlow / hrhigh | 50 / 120 | Heart-rate limits (bpm) |
| spo2low | 92 | SpO₂ lower limit (%) |
| temphigh | 38 | Temperature upper limit (°C) |
| ssid / pass | — | Wi-Fi network (set with the wifi: command) |
| pin | 1234 | BLE PIN |

Table 11: Settings stored in flash (change with set:key=value)

# 7. Setup and Deployment Guide

## 7.1 Hardware Assembly

- Place the ESP32 DevKit across the centre gap of the breadboard so both pin rows are free.

- Connect the ESP32 3V3 pin to one breadboard power rail (3.3 V), VIN to the other rail (5 V) and both GND pins to the ground rails.

- Wire each module exactly as in the pin connection table (section 4.2), including the resistor dividers marked with [n] in the circuit diagram.

- Check every connection against the circuit diagram before powering on. Make sure no 5 V signal reaches an ESP32 pin directly.

- Connect the USB cable to a computer (for flashing) or a 5 V adapter (for normal use).

## 7.2 Flashing the Firmware

### Option A — Arduino IDE

- Install Arduino IDE 2.x from arduino.cc.

- File → Preferences → Additional boards manager URLs: https://espressif.github.io/arduino-esp32/package_esp32_index.json.

- Tools → Board → Boards Manager: install esp32 by Espressif Systems (version 3.x).

- Sketch → Include Library → Manage Libraries: install NimBLE-Arduino (2.x), ArduinoJson (7.x), PubSubClient, SparkFun MAX3010x Pulse and Proximity Sensor Library, OneWire, DallasTemperature, Adafruit SSD1306, Adafruit GFX Library.

- Download the repository and open the folder P08_Health_Monitoring_System (it already contains IoTCore.h, Cloud.h and config.h).

- Tools: Board ESP32 Dev Module, Partition Scheme Minimal SPIFFS (1.9MB APP with OTA/190KB SPIFFS), Upload Speed 921600, and the correct COM port.

- Click Upload. If it stays at 'Connecting…', hold the BOOT button on the ESP32 until upload starts.

- Open Serial Monitor at 115200 baud and type status to check the device.

### Option B — no software installation (GitHub Actions + browser flasher)

- Open the repository https://github.com/rohitkukreja705/esp32-iot-projects → Actions → latest Build firmware run.

- Download the artifact firmware-P08_Health_Monitoring_System and unzip it to get P08_Health_Monitoring_System.ino.merged.bin.

- Open https://espressif.github.io/esptool-js/ in Chrome or Edge on a computer, click Connect and select the ESP32's serial port.

- Set Flash Address 0x0, choose the merged .bin file and click Program. Press the EN (reset) button when done.

## 7.3 Cloud Setup (ThingSpeak)

- Go to thingspeak.mathworks.com and sign up with a free MathWorks account (choose the free, non-commercial licence).

- Open Channels → My Channels → New Channel. Name it 'Health Monitoring System'.

- Tick and name the fields exactly as listed in the table below, then click Save Channel.

- Open the API Keys tab and copy the Write API Key (16 characters).

- In the BLE web console open Setup, unlock with the PIN and save the key in the tskey field (equivalent command: set:tskey=YOURKEY).

- Open Private View to see live charts. Add Gauge or Numeric Display widgets with Add Widgets, and use Sharing → Share channel view with everyone if you want a public link.

| Field / feed / table | Content |
| --- | --- |
| Field 1 | Heart rate (bpm) |
| Field 2 | SpO₂ (%) |
| Field 3 | Temperature (°C) |
| Field 4 | Alert flag (0/1) |
| Status | Alert text or 'normal' |

Table 12: Items to create in ThingSpeak

## 7.4 BLE App Setup (Web Bluetooth Console)

No app installation is needed. The console is a web page hosted free on GitHub Pages; Chrome talks to the ESP32 directly through the Web Bluetooth API.

- On an Android phone turn on Bluetooth and Location (Android requires location for BLE scanning).

- Open https://rohitkukreja705.github.io/esp32-iot-projects/?p=P08 in Chrome. (On iPhone use the free Bluefy browser; on a laptop use Chrome or Edge.)

- Tap Connect and choose IOT-P08-XXXX in the pop-up list. The dot turns green and the Live tab starts updating.

- Open Setup, enter the PIN (default 1234) and tap Unlock.

- Enter your 2.4 GHz Wi-Fi name and password and tap Save Wi-Fi. After a few seconds the status line shows the IP address.

- Enter the ThingSpeak keys from section 7.3 and tap Save next to each.

- Use the Control tab for commands and the Log tab to see every event and acknowledgement.

### Alternative: nRF Connect app

- Install nRF Connect for Mobile (Nordic Semiconductor, free) and scan.

- Connect to IOT-P08-XXXX and open Nordic UART Service.

- Enable notifications on the TX characteristic (6E400003…).

- Write UTF-8 text to the RX characteristic (6E400002…), e.g. auth:1234, then status.

To host the console yourself: in the GitHub repository open Settings → Pages, choose 'Deploy from a branch', branch main, folder /docs. Web Bluetooth only works on HTTPS pages, which GitHub Pages provides.

## 7.5 First-Run Checklist

- Serial Monitor shows '[BLE] advertising as IOT-P08-XXXX'.

- Status in the console shows wifi: true and an IP address.

- Data appears in ThingSpeak within a minute.

- All sensor values in the Live tab are plausible (no '—').

# 8. Testing and Results

The following test cases verify each function. Record the observed result and mark Pass/Fail while testing your build.

| # | Test | Procedure | Expected result | Observed |
| --- | --- | --- | --- | --- |
| T1 | Sensor detection | Boot with Serial Monitor open | No '[MAX30102] not found' message |  |
| T2 | Finger detection | Place and remove a finger | 'Finger detected' / 'Finger removed' events |  |
| T3 | Heart rate | Rest finger 15 s; compare with a commercial oximeter or manual pulse count | Within ±5 bpm |  |
| T4 | SpO₂ | Compare with a commercial finger oximeter at rest | Within ±2-3 % (typical 95-99 %) |  |
| T5 | Temperature | Hold probe under the arm for 3 minutes; compare with a clinical thermometer | Within ±0.5 °C |  |
| T6 | Alert | Set hrhigh=60 over BLE | Buzzer chirps and 'vital_alert' event appears |  |
| T7 | Cloud | Keep finger on sensor for 1 minute | ThingSpeak fields 1-4 update every 20 s |  |

Table 13: Test cases

## 8.1 Expected Output

A typical data frame received in the console's Log tab looks like this (values will differ):

{"hr":76,"spo2":98,"temp":36.6,"finger":true,"alert":false,"type":"data"}

Serial Monitor (115200 baud) prints the same lines along with connection messages such as [WiFi] connected, [BLE] connected and the HTTP or MQTT result of each cloud upload, which helps in troubleshooting.

## 8.2 Troubleshooting

| Symptom | Likely cause and fix |
| --- | --- |
| Device not in the Connect list | Bluetooth/Location off on the phone, or the page is not opened in Chrome over HTTPS. |
| 'Sketch too big' on upload | Select Partition Scheme 'Minimal SPIFFS (1.9MB APP with OTA)'. |
| Wi-Fi never connects | ESP32 supports only 2.4 GHz networks; check the SSID/password (case-sensitive). |
| Settings refused | Send auth:<PIN> first; the unlock is per BLE connection. |
| Sensor shows null / — | Check wiring and supply voltage; see the Serial Monitor for 'not found' messages. |
| Board resets randomly | Weak USB supply or servo/relay current; use a 5 V 2 A adapter and add a 470 µF capacitor. |
| No data in ThingSpeak | Wrong Write API Key or interval below 15 s. |

Table 14: Troubleshooting guide

# 9. Discussion

## 9.1 Advantages

- Three vitals in one low-cost device with remote logging.

- Responsive: non-blocking design keeps BLE and Wi-Fi working while measuring.

- Uploads only valid data (finger present).

- Limits are configurable per patient.

## 9.2 Limitations

- Not a medical device; accuracy depends on finger pressure, motion, ambient light and skin tone.

- Maxim's reference algorithm is sensitive to motion; readings at rest only.

- FIFO samples may be lost during the 1-2 s HTTPS upload, briefly delaying the next estimate.

- ThingSpeak data should not hold identifiable patient information.

## 9.3 Safety Precautions

- Educational prototype only — never use it to diagnose, treat or replace professional medical equipment.

- Use a USB power bank or isolated adapter; never connect a person to a device powered from a non-isolated supply.

- Clean the sensor and probe between users.

## 9.4 Security Considerations

- Secrets (Wi-Fi password, cloud keys) are stored in the ESP32's NVS flash, not in source code.

- BLE commands that change the device require a PIN; change the default PIN after setup.

- Cloud traffic is TLS-encrypted. Certificate verification is disabled (setInsecure) for simplicity; for production, pin the server's root certificate with setCACert().

- Use a dedicated cloud account/project for the device and rotate keys if the device is lost.

## 9.5 Future Scope

- Use a better algorithm (bandpass filter + peak detection, or Maxim's SpO₂ calibration) and a finger clip.

- Add ECG (AD8232) and blood pressure estimation from PTT.

- Send SMS/email to caregivers via ThingSpeak React + ThingHTTP.

- Wearable version with a LiPo battery and wrist-worn sensor.

# 10. Conclusion

The Health Monitoring System demonstrates how optical PPG and digital temperature sensing can be combined on an ESP32 into a connected vital-signs monitor. Its non-blocking design keeps BLE and Wi-Fi responsive, finger detection keeps cloud data clean, and configurable alerts make it a useful learning platform for biomedical IoT.

All design goals listed in Chapter 1 are met with low-cost parts available in the Indian market, free cloud services and open-source software. The shared firmware core and the single Web Bluetooth console make the project easy to reproduce and extend.

# References

- Maxim Integrated (Analog Devices), MAX30102 datasheet and reference design (RD117) algorithm.

- SparkFun Electronics, SparkFun MAX3010x Sensor Library, GitHub.

- Maxim Integrated, DS18B20 Programmable Resolution 1-Wire Digital Thermometer datasheet.

- J. Allen (2007), Photoplethysmography and its application in clinical physiological measurement, Physiol. Meas. 28.

- Espressif Systems, ESP32 Series Datasheet and ESP32 Technical Reference Manual, espressif.com.

- Espressif Systems, Arduino-ESP32 documentation, docs.espressif.com/projects/arduino-esp32.

- H2zero, NimBLE-Arduino library documentation, github.com/h2zero/NimBLE-Arduino.

- Bluetooth SIG, Bluetooth Core Specification 4.2 — Generic Attribute Profile (GATT).

- W3C Web Bluetooth Community Group, Web Bluetooth specification, webbluetoothcg.github.io/web-bluetooth.

- B. Blanchon, ArduinoJson 7 documentation, arduinojson.org.

- MathWorks, ThingSpeak documentation — Write data with REST API, mathworks.com/help/thingspeak.

# Appendix A: Project Source Code

The complete code is in the repository folder P08_Health_Monitoring_System (https://github.com/rohitkukreja705/esp32-iot-projects).

config.h

#pragma once

// ===== P08 Health Monitoring System — configuration =====

#define WIFI_SSID     ""

#define WIFI_PASS     ""

#define DEVICE_PIN    "1234"

#define TS_WRITE_KEY  ""

// ----- pins -----

// MAX30102 + SSD1306 share I2C: SDA 21, SCL 22 (MAX30102 0x57, OLED 0x3C)

#define PIN_DS18B20   4      // DS18B20 data + 4.7k pull-up to 3.3 V

#define PIN_BUZZER    18

#define OLED_ADDR     0x3C

// ----- alert thresholds (set:hrlow= set:hrhigh= set:spo2low= set:templow/temphigh=) -----

#define DEF_HR_LOW     50

#define DEF_HR_HIGH    120

#define DEF_SPO2_LOW   92

#define DEF_TEMP_HIGH  38.0

#define FINGER_IR_MIN  50000  // IR level that means a finger is on the sensor

#define TS_INTERVAL_S  20

P08_Health_Monitoring_System.ino

/*

* P08 — IoT Health Monitoring System (Heart Rate, SpO2, Body Temperature)

* =======================================================================

* Peripherals: MAX30102 pulse-oximeter, DS18B20 temperature probe,

*              SSD1306 OLED, buzzer

* Wi-Fi  : ThingSpeak — HR, SpO2, temperature, alert flag every 20 s

*          (only while a finger is on the sensor and readings are valid)

* BLE    : live vitals on the phone, alert events, configurable limits

* Extras : non-blocking 25 Hz sampling with a sliding 4-second window,

*          Maxim SpO2/HR algorithm, finger detection, averaged HR,

*          non-blocking DS18B20 conversion

*

* DISCLAIMER: educational project, NOT a medical device. Do not use it for

* diagnosis or treatment decisions.

*/

#include "config.h"

#include "IoTCore.h"

#include "Cloud.h"

#include <Wire.h>

#include <MAX30105.h>

#include <spo2_algorithm.h>

#include <OneWire.h>

#include <DallasTemperature.h>

#include <Adafruit_GFX.h>

#include <Adafruit_SSD1306.h>

MAX30105 pox;

OneWire oneWire(PIN_DS18B20);

DallasTemperature ds(&oneWire);

Adafruit_SSD1306 oled(128, 64, &Wire, -1);

bool poxOK = false, oledOK = false;

const int BUF = 100;                    // 4 s at 25 samples/s

uint32_t irBuf[BUF], redBuf[BUF];

int filled = 0, newSamples = 0;

bool finger = false;

float hr = NAN, spo2 = NAN, bodyTemp = NAN;

float hrHist[4]; int hrN = 0, hrI = 0;

bool alertOn = false; String alertMsg = "";

uint32_t lastTempReq = 0, lastUpload = 0, lastUi = 0;

float hrLow, hrHigh, spo2Low, tempHigh;

void loadSettings() {

hrLow = iot::settingF("hrlow", DEF_HR_LOW);   hrHigh = iot::settingF("hrhigh", DEF_HR_HIGH);

spo2Low = iot::settingF("spo2low", DEF_SPO2_LOW); tempHigh = iot::settingF("temphigh", DEF_TEMP_HIGH);

}

void processWindow() {

int32_t s, h; int8_t sv, hv;

maxim_heart_rate_and_oxygen_saturation(irBuf, BUF, redBuf, &s, &sv, &h, &hv);

if (hv && h > 30 && h < 220) {

hrHist[hrI] = h; hrI = (hrI + 1) % 4; if (hrN < 4) hrN++;

float sum = 0; for (int i = 0; i < hrN; i++) sum += hrHist[i];

hr = sum / hrN;

}

if (sv && s > 70 && s <= 100) spo2 = s;

}

void readPox() {

if (!poxOK) return;

pox.check();

while (pox.available()) {

uint32_t red = pox.getFIFORed(), ir = pox.getFIFOIR();

pox.nextSample();

bool f = ir > FINGER_IR_MIN;

if (f != finger) {

finger = f; filled = 0; newSamples = 0; hrN = 0;

if (!f) { hr = NAN; spo2 = NAN; }

iot::sendEvent("finger", f ? "Finger detected - hold still" : "Finger removed");

}

if (!finger) continue;

if (filled < BUF) { irBuf[filled] = ir; redBuf[filled] = red; filled++; }

else {                               // slide the window by one sample

memmove(irBuf, irBuf + 1, (BUF - 1) * sizeof(uint32_t));

memmove(redBuf, redBuf + 1, (BUF - 1) * sizeof(uint32_t));

irBuf[BUF - 1] = ir; redBuf[BUF - 1] = red;

}

if (filled == BUF && ++newSamples >= 25) { newSamples = 0; processWindow(); }   // every 1 s

}

}

void readTemp() {

if (millis() - lastTempReq < 1000) return;

float t = ds.getTempCByIndex(0);

if (t > -50 && t < 85) bodyTemp = t;   // -127 = sensor missing

ds.requestTemperatures();              // non-blocking; read on next call

lastTempReq = millis();

}

void checkAlerts() {

String m = "";

if (finger && !isnan(hr) && (hr < hrLow || hr > hrHigh)) m += "HR " + String(hr, 0) + " bpm. ";

if (finger && !isnan(spo2) && spo2 < spo2Low) m += "SpO2 " + String(spo2, 0) + "%. ";

if (!isnan(bodyTemp) && bodyTemp > tempHigh) m += "Temp " + String(bodyTemp, 1) + "C. ";

bool a = m.length() > 0;

if (a && (!alertOn || m != alertMsg)) iot::sendEvent("vital_alert", m);

alertOn = a; alertMsg = m;

digitalWrite(PIN_BUZZER, alertOn && (millis() / 200) % 5 == 0);

}

void drawOled() {

if (!oledOK) return;

oled.clearDisplay(); oled.setTextColor(SSD1306_WHITE);

oled.setTextSize(1); oled.setCursor(0, 0);

oled.printf("HEALTH %s %s", iot::wifiOK() ? "WiFi" : "----", iot::bleConnected() ? "BLE" : "");

if (!finger) { oled.setCursor(0, 24); oled.setTextSize(2); oled.print("Place"); oled.setCursor(0, 42); oled.print("finger"); }

else {

oled.setTextSize(2); oled.setCursor(0, 14);

oled.print(isnan(hr) ? String("--") : String(hr, 0)); oled.setTextSize(1); oled.print(" bpm");

if ((millis() / 400) % 2) oled.fillCircle(118, 20, 4, SSD1306_WHITE);   // heartbeat dot

oled.setTextSize(2); oled.setCursor(0, 34);

oled.print(isnan(spo2) ? String("--") : String(spo2, 0)); oled.setTextSize(1); oled.print(" %SpO2");

}

oled.setTextSize(1); oled.setCursor(0, 56);

oled.print(isnan(bodyTemp) ? String("Temp --") : "Temp " + String(bodyTemp, 1) + "C");

if (alertOn) oled.print("  ALERT");

oled.display();

}

void sendBle() {

JsonDocument d;

d["hr"] = iot::rnd(hr, 0); d["spo2"] = iot::rnd(spo2, 0); d["temp"] = iot::rnd(bodyTemp);

d["finger"] = finger; d["alert"] = alertOn;

if (alertOn) d["msg"] = alertMsg;

iot::sendData(d);

}

bool onCommand(const String& cmd, const String& arg, bool authed) {

if (cmd == "setting") { loadSettings(); return true; }

return false;

}

void setup() {

pinMode(PIN_BUZZER, OUTPUT);

iot::begin("P08", onCommand);

loadSettings();

Wire.begin(21, 22);

poxOK = pox.begin(Wire, I2C_SPEED_FAST);

if (poxOK) pox.setup(60, 4, 2, 100, 411, 4096);   // brightness, avg, Red+IR, 100 sps, 411 us, range

else Serial.println("[MAX30102] not found");

ds.begin(); ds.setWaitForConversion(false); ds.requestTemperatures();

oledOK = oled.begin(SSD1306_SWITCHCAPVCC, OLED_ADDR);

}

void loop() {

iot::loop();

readPox();

readTemp();

checkAlerts();

uint32_t now = millis();

if (now - lastUi >= 1000) { lastUi = now; drawOled(); sendBle(); }

if (iot::wifiOK() && finger && !isnan(hr) && now - lastUpload >= TS_INTERVAL_S * 1000UL) {

lastUpload = now;

float f[4] = { hr, spo2, bodyTemp, alertOn ? 1.0f : 0.0f };

cloud::thingSpeakUpdate(f, 4, alertOn ? alertMsg : "normal");

}

}

# Appendix B: Shared Library Code

These two headers are identical in every project folder (master copies in /common).

IoTCore.h

/*

* IoTCore.h  —  shared runtime for all ESP32 IoT projects in this repo

* -------------------------------------------------------------------

*  - Wi-Fi station with automatic reconnect (credentials stored in flash)

*  - BLE server using the Nordic UART Service (NUS) layout, so the device

*    works with the Web Bluetooth dashboard in /docs AND with generic apps

*    such as "nRF Connect" or "Serial Bluetooth Terminal"

*  - Line-based command protocol shared by BLE and USB Serial

*  - Persistent settings (Preferences / NVS): Wi-Fi, cloud keys, PIN, thresholds

*  - NTP time sync (IST) for timestamps

*

*  Wire protocol

*  -------------

*  Phone -> ESP32 (write to RX characteristic, text):

*      auth:1234              unlock control + settings for this connection

*      status                 device replies with a status JSON line

*      set:key=value          store a setting (Wi-Fi, cloud key, threshold…)

*      get:key                read back a setting (secrets are masked)

*      wifi:SSID,PASSWORD     shortcut for set:ssid + set:pass + reconnect

*      reboot                 restart the board

*      <anything else>        forwarded to the project's own command handler

*  ESP32 -> Phone (notify on TX characteristic):

*      one JSON object per line, terminated by '\n'

*      {"type":"data",...}  {"type":"status",...}  {"type":"ack",...}

*      {"type":"event",...} {"type":"err",...}

*  Lines longer than one BLE packet are split into chunks; the receiver

*  joins chunks until it sees '\n'.

*

*  Library: NimBLE-Arduino 2.x (lighter than Bluedroid, leaves more RAM for TLS)

*  Board:   ESP32 Dev Module, Partition Scheme "Minimal SPIFFS (1.9MB APP with OTA)"

*/

#pragma once

#include <Arduino.h>

#include <WiFi.h>

#include <Preferences.h>

#include <NimBLEDevice.h>

#include <ArduinoJson.h>

#include <time.h>

#ifndef DEVICE_PIN

#define DEVICE_PIN "1234"

#endif

#ifndef WIFI_SSID

#define WIFI_SSID ""

#endif

#ifndef WIFI_PASS

#define WIFI_PASS ""

#endif

namespace iot {

// Nordic UART Service UUIDs (widely supported by BLE terminal apps)

static const char* NUS_SERVICE = "6E400001-B5A3-F393-E0A9-E50E24DCCA9E";

static const char* NUS_RX      = "6E400002-B5A3-F393-E0A9-E50E24DCCA9E";  // phone writes here

static const char* NUS_TX      = "6E400003-B5A3-F393-E0A9-E50E24DCCA9E";  // ESP32 notifies here

// Project command handler: return true if the command was recognised.

using CommandHandler = bool (*)(const String& cmd, const String& arg, bool authed);

// ------------------------------------------------------------------ state

struct CmdMsg { char text[256]; bool fromBle; };

static Preferences        prefs;

static NimBLEServer*      bleServer   = nullptr;

static NimBLECharacteristic* txChr    = nullptr;

static QueueHandle_t      cmdQueue    = nullptr;

static CommandHandler     userHandler = nullptr;

static String             deviceName;

static String             projectId;

static volatile uint16_t  bleMtu      = 23;

static volatile bool      bleAuthed   = false;   // reset on every new connection

static bool               serialAuthed = true;   // USB = physical access, trusted

static bool               timeSynced  = false;

static uint32_t           lastWifiTry = 0;

static String             rxLine;                // assembles multi-packet writes

// --------------------------------------------------------------- settings

inline String setting(const char* key, const char* def = "") {

return prefs.isKey(key) ? prefs.getString(key, def) : String(def);

}

inline void saveSetting(const char* key, const String& v) { prefs.putString(key, v); }

inline float settingF(const char* key, float def) {

String s = setting(key, "");

return s.length() ? s.toFloat() : def;

}

// ------------------------------------------------------------------- BLE

inline bool bleConnected() { return bleServer && bleServer->getConnectedCount() > 0; }

// Send one line (adds '\n'), split to fit the negotiated MTU.

inline void bleSend(const String& line) {

if (!bleConnected() || !txChr) return;

String msg = line + "\n";

size_t chunk = (bleMtu > 23 ? bleMtu : 23) - 3;

for (size_t i = 0; i < msg.length(); i += chunk) {

size_t n = min(chunk, msg.length() - i);

txChr->setValue((const uint8_t*)msg.c_str() + i, n);

txChr->notify();

if (msg.length() > chunk) delay(8);   // give the stack time between chunks

}

}

// Send to BLE and mirror to USB Serial.

inline void emit(const String& line) {

Serial.println(line);

bleSend(line);

}

inline void emitAck(const String& cmd, bool ok, const String& msg = "") {

JsonDocument d;

d["type"] = ok ? "ack" : "err";

d["cmd"] = cmd;

if (msg.length()) d["msg"] = msg;

String s; serializeJson(d, s); emit(s);

}

inline void enqueue(const String& s, bool fromBle) {

if (!cmdQueue || s.length() == 0) return;

CmdMsg m; memset(&m, 0, sizeof(m));

strncpy(m.text, s.c_str(), sizeof(m.text) - 1);

m.fromBle = fromBle;

xQueueSend(cmdQueue, &m, 0);

}

class ServerCB : public NimBLEServerCallbacks {

void onConnect(NimBLEServer* s, NimBLEConnInfo& info) override {

bleAuthed = false;

rxLine = "";

Serial.printf("[BLE] connected %s\n", info.getAddress().toString().c_str());

}

void onDisconnect(NimBLEServer* s, NimBLEConnInfo& info, int reason) override {

bleAuthed = false;

bleMtu = 23;

Serial.printf("[BLE] disconnected (reason %d), advertising again\n", reason);

// NimBLE restarts advertising automatically (advertiseOnDisconnect = true)

}

void onMTUChange(uint16_t mtu, NimBLEConnInfo& info) override {

bleMtu = mtu;

Serial.printf("[BLE] MTU %u\n", mtu);

}

};

class RxCB : public NimBLECharacteristicCallbacks {

void onWrite(NimBLECharacteristic* c, NimBLEConnInfo& info) override {

std::string v = c->getValue();

// Accept either a single write without newline, or several writes

// that together form lines terminated by '\n'.

for (char ch : v) {

if (ch == '\r') continue;

if (ch == '\n') { enqueue(rxLine, true); rxLine = ""; }

else if (rxLine.length() < 250) rxLine += ch;

}

if (v.find('\n') == std::string::npos && rxLine.length() && rxLine.length() == v.length()) {

enqueue(rxLine, true);  // plain one-shot write (e.g. typed in nRF Connect)

rxLine = "";

}

}

};

inline void bleBegin() {

uint8_t mac[6]; WiFi.macAddress(mac);

char suffix[6]; snprintf(suffix, sizeof(suffix), "%02X%02X", mac[4], mac[5]);

deviceName = "IOT-" + projectId + "-" + suffix;

NimBLEDevice::init(deviceName.c_str());

NimBLEDevice::setMTU(247);

bleServer = NimBLEDevice::createServer();

bleServer->setCallbacks(new ServerCB());

NimBLEService* svc = bleServer->createService(NUS_SERVICE);

txChr = svc->createCharacteristic(NUS_TX, NIMBLE_PROPERTY::NOTIFY | NIMBLE_PROPERTY::READ);

NimBLECharacteristic* rx = svc->createCharacteristic(

NUS_RX, NIMBLE_PROPERTY::WRITE | NIMBLE_PROPERTY::WRITE_NR, 512);

rx->setCallbacks(new RxCB());

// (NimBLE 2.x starts services automatically when advertising begins)

// 128-bit UUID goes in the advertisement, the name in the scan response

// (both together would not fit in one 31-byte packet).

NimBLEAdvertisementData adv, scan;

adv.setFlags(BLE_HS_ADV_F_DISC_GEN | BLE_HS_ADV_F_BREDR_UNSUP);

adv.setCompleteServices(NimBLEUUID(NUS_SERVICE));

scan.setName(deviceName.c_str());

NimBLEAdvertising* a = NimBLEDevice::getAdvertising();

a->setAdvertisementData(adv);

a->setScanResponseData(scan);

a->start();

Serial.printf("[BLE] advertising as %s\n", deviceName.c_str());

}

// ------------------------------------------------------------------ Wi-Fi

inline bool wifiOK() { return WiFi.status() == WL_CONNECTED; }

inline void wifiConnect() {

String ssid = setting("ssid", WIFI_SSID);

String pass = setting("pass", WIFI_PASS);

if (ssid.isEmpty()) {

Serial.println("[WiFi] no SSID yet - send  wifi:SSID,PASSWORD  over BLE");

return;

}

Serial.printf("[WiFi] connecting to %s\n", ssid.c_str());

WiFi.disconnect();

WiFi.begin(ssid.c_str(), pass.c_str());

lastWifiTry = millis();

}

inline void wifiLoop() {

static bool wasUp = false;

bool up = wifiOK();

if (up && !wasUp) {

Serial.printf("[WiFi] connected, IP %s, RSSI %d\n",

WiFi.localIP().toString().c_str(), WiFi.RSSI());

configTzTime("IST-5:30", "pool.ntp.org", "time.google.com");

}

if (!up && millis() - lastWifiTry > 20000) wifiConnect();   // retry every 20 s

wasUp = up;

if (up && !timeSynced && time(nullptr) > 1700000000) {

timeSynced = true;

Serial.println("[NTP] time synced");

}

}

// ISO-8601 local time, or uptime if NTP has not synced yet.

inline String timestamp() {

time_t now = time(nullptr);

if (now < 1700000000) return "uptime+" + String(millis() / 1000) + "s";

struct tm t; localtime_r(&now, &t);

char buf[32]; strftime(buf, sizeof(buf), "%Y-%m-%dT%H:%M:%S+05:30", &t);

return String(buf);

}

// ------------------------------------------------------- built-in commands

inline bool isSecretKey(const String& k) {

return k == "pass" || k.endsWith("key") || k == "pin";

}

inline void sendStatus() {

JsonDocument d;

d["type"] = "status";

d["dev"] = deviceName;

d["wifi"] = wifiOK();

d["ssid"] = setting("ssid", WIFI_SSID);

d["rssi"] = wifiOK() ? WiFi.RSSI() : 0;

d["ip"] = wifiOK() ? WiFi.localIP().toString() : "";

d["auth"] = bleAuthed;

d["heap"] = ESP.getFreeHeap();

d["up"] = millis() / 1000;

d["time"] = timestamp();

String s; serializeJson(d, s); emit(s);

}

inline void handleLine(const String& raw, bool fromBle) {

String line = raw; line.trim();

if (line.isEmpty()) return;

int c = line.indexOf(':');

String cmd = (c < 0 ? line : line.substring(0, c));

String arg = (c < 0 ? "" : line.substring(c + 1));

cmd.trim(); cmd.toLowerCase();

bool authed = fromBle ? bleAuthed : serialAuthed;

if (cmd == "auth") {

bool ok = arg == setting("pin", DEVICE_PIN);

if (fromBle) bleAuthed = ok;

emitAck(cmd, ok, ok ? "unlocked" : "wrong PIN");

return;

}

if (cmd == "status") { sendStatus(); return; }

if (cmd == "ping")   { emitAck(cmd, true, "pong"); return; }

// Everything below changes the device, so it needs the PIN over BLE.

if (cmd == "set" || cmd == "wifi" || cmd == "reboot" || cmd == "get" || cmd == "forget") {

if (!authed) { emitAck(cmd, false, "send auth:<PIN> first"); return; }

if (cmd == "set") {

int e = arg.indexOf('=');

if (e <= 0) { emitAck(cmd, false, "use set:key=value"); return; }

String k = arg.substring(0, e); k.trim();

String v = arg.substring(e + 1); v.trim();

if (k.length() > 15) { emitAck(cmd, false, "key too long"); return; }

saveSetting(k.c_str(), v);

emitAck(cmd, true, k + " saved");

if (k == "ssid" || k == "pass") wifiConnect();

// let the project react to changed thresholds etc.

if (userHandler) userHandler("setting", k, true);

return;

}

if (cmd == "get") {

String v = setting(arg.c_str(), "");

if (isSecretKey(arg) && v.length()) v = "****" + v.substring(max(0, (int)v.length() - 3));

emitAck(cmd, true, arg + "=" + v);

return;

}

if (cmd == "wifi") {

int k = arg.indexOf(',');

String ssid = k < 0 ? arg : arg.substring(0, k);

String pass = k < 0 ? "" : arg.substring(k + 1);

saveSetting("ssid", ssid); saveSetting("pass", pass);

emitAck(cmd, true, "connecting to " + ssid);

wifiConnect();

return;

}

if (cmd == "forget") { prefs.clear(); emitAck(cmd, true, "all settings erased"); return; }

if (cmd == "reboot") { emitAck(cmd, true, "rebooting"); delay(300); ESP.restart(); }

}

if (userHandler && userHandler(cmd, arg, authed)) return;

emitAck(cmd, false, "unknown command");

}

// ------------------------------------------------------------ entry points

// Call once from setup().  id = short project id, e.g. "P01".

inline void begin(const char* id, CommandHandler handler) {

Serial.begin(115200);

delay(200);

projectId = id;

userHandler = handler;

prefs.begin("iot", false);

cmdQueue = xQueueCreate(8, sizeof(CmdMsg));

WiFi.mode(WIFI_STA);

WiFi.setAutoReconnect(true);

// NOTE: Wi-Fi modem sleep must stay ON while BLE is running (ESP32 coexistence).

bleBegin();

wifiConnect();

Serial.println("[CORE] ready - type 'status' here or connect over BLE");

}

// Call every loop() iteration.

inline void loop() {

wifiLoop();

// USB serial commands (handy while developing)

static String serialLine;

while (Serial.available()) {

char ch = Serial.read();

if (ch == '\n' || ch == '\r') { if (serialLine.length()) enqueue(serialLine, false); serialLine = ""; }

else if (serialLine.length() < 250) serialLine += ch;

}

CmdMsg m;

while (cmdQueue && xQueueReceive(cmdQueue, &m, 0) == pdTRUE) handleLine(String(m.text), m.fromBle);

}

// Round for compact JSON. NaN stays NaN, which ArduinoJson writes as null.

inline float rnd(float v, int dp = 1) {

if (isnan(v)) return v;

float m = powf(10, dp);

return roundf(v * m) / m;

}

// Helper: build and send a {"type":"data",...} line from a filled document.

inline void sendData(JsonDocument& d) {

d["type"] = "data";

String s; serializeJson(d, s); emit(s);

}

inline void sendEvent(const String& name, const String& msg) {

JsonDocument d;

d["type"] = "event"; d["name"] = name; d["msg"] = msg; d["time"] = timestamp();

String s; serializeJson(d, s); emit(s);

}

}  // namespace iot

Cloud.h

/*

* Cloud.h  —  free-tier cloud helpers used by the projects

* ---------------------------------------------------------

*  ThingSpeak   : HTTPS REST "update" (free tier: 1 update / 15 s per channel)

*  Adafruit IO  : MQTT over TLS (free tier: 30 data points / minute, 10 feeds)

*  Supabase     : PostgREST over HTTPS (free tier: 500 MB Postgres)

*

*  Keys are read from settings (stored over BLE with set:key=value) so no

*  secret has to be hard-coded or committed to GitHub:

*     tskey    ThingSpeak Write API Key

*     aiouser  Adafruit IO username        aiokey  Adafruit IO key

*     sburl    https://<ref>.supabase.co   sbkey   Supabase publishable/anon key

*

*  TLS note: certificate checking is disabled with setInsecure() to keep the

*  projects beginner-friendly and independent of CA rotations. Traffic is

*  still encrypted. For production, pin the root CA with setCACert().

*/

#pragma once

#include <WiFiClientSecure.h>

#include <HTTPClient.h>

#include <PubSubClient.h>

#include "IoTCore.h"

// Optional compile-time defaults (normally left empty and set over BLE)

#ifndef TS_WRITE_KEY

#define TS_WRITE_KEY ""

#endif

#ifndef AIO_USERNAME

#define AIO_USERNAME ""

#endif

#ifndef AIO_KEY

#define AIO_KEY ""

#endif

#ifndef SUPABASE_URL

#define SUPABASE_URL ""

#endif

#ifndef SUPABASE_KEY

#define SUPABASE_KEY ""

#endif

namespace cloud {

// ============================================================== ThingSpeak

// vals[i] -> field(i+1). Pass NAN to skip a field. Returns true on success.

inline bool thingSpeakUpdate(const float* vals, int n, const String& status = "") {

if (!iot::wifiOK()) return false;

String key = iot::setting("tskey", TS_WRITE_KEY);

if (key.isEmpty()) { Serial.println("[TS] no tskey set"); return false; }

String url = "https://api.thingspeak.com/update?api_key=" + key;

for (int i = 0; i < n && i < 8; i++)

if (!isnan(vals[i])) url += "&field" + String(i + 1) + "=" + String(vals[i], 2);

if (status.length()) {

String st = status; st.replace(" ", "%20");

url += "&status=" + st;

}

WiFiClientSecure client; client.setInsecure();

HTTPClient http; http.setTimeout(8000);

if (!http.begin(client, url)) return false;

int code = http.GET();

String body = http.getString();

http.end();

bool ok = code == 200 && body.toInt() > 0;   // ThingSpeak returns the entry id, 0 = rejected

Serial.printf("[TS] HTTP %d entry=%s\n", code, body.c_str());

return ok;

}

// ================================================================ Supabase

// Generic PostgREST call. method: GET/POST/PATCH. path e.g. "/rest/v1/readings"

inline int supabaseRequest(const char* method, const String& path, const String& body,

String* response = nullptr, const char* prefer = "return=minimal") {

if (!iot::wifiOK()) return -1;

String base = iot::setting("sburl", SUPABASE_URL);

String key  = iot::setting("sbkey", SUPABASE_KEY);

if (base.isEmpty() || key.isEmpty()) { Serial.println("[SB] sburl/sbkey not set"); return -2; }

if (base.endsWith("/")) base.remove(base.length() - 1);

WiFiClientSecure client; client.setInsecure();

HTTPClient http; http.setTimeout(8000);

if (!http.begin(client, base + path)) return -3;

http.addHeader("apikey", key);

http.addHeader("Authorization", "Bearer " + key);

http.addHeader("Content-Type", "application/json");

if (prefer) http.addHeader("Prefer", prefer);

int code = http.sendRequest(method, body);

if (response) *response = http.getString();

else if (code >= 300) Serial.printf("[SB] %s\n", http.getString().c_str());

http.end();

Serial.printf("[SB] %s %s -> %d\n", method, path.c_str(), code);

return code;

}

inline bool supabaseInsert(const char* table, const String& json) {

int c = supabaseRequest("POST", String("/rest/v1/") + table, json);

return c >= 200 && c < 300;

}

// filter is a PostgREST filter, e.g. "id=eq.1"

inline bool supabaseUpdate(const char* table, const String& filter, const String& json) {

int c = supabaseRequest("PATCH", String("/rest/v1/") + table + "?" + filter, json);

return c >= 200 && c < 300;

}

// query e.g. "select=*&id=eq.1"  -> JSON array text in out

inline bool supabaseSelect(const char* table, const String& query, String& out) {

int c = supabaseRequest("GET", String("/rest/v1/") + table + "?" + query, "", &out, nullptr);

return c == 200;

}

// ============================================================ Adafruit IO

// MQTT wrapper with a simple rate limiter to respect the 30 msg/min limit.

using FeedCallback = void (*)(const String& feed, const String& value);

class AdafruitIO {

public:

// requestLast = true asks Adafruit IO for each feed's last value on connect

// (useful for switches changed while offline; unsafe for things like pumps)

void begin(FeedCallback cb, const char* const* subFeeds, int nSub, bool requestLast = true) {

cb_ = cb; subs_ = subFeeds; nSub_ = nSub; requestLast_ = requestLast;

tls_.setInsecure();

mqtt_.setClient(tls_);

mqtt_.setServer("io.adafruit.com", 8883);

mqtt_.setBufferSize(512);

mqtt_.setKeepAlive(60);

mqtt_.setCallback([](char* topic, byte* payload, unsigned int len) {

String t(topic), v;

for (unsigned int i = 0; i < len; i++) v += (char)payload[i];

int p = t.lastIndexOf('/');

String feed = t.substring(p + 1);

if (instance()->cb_) instance()->cb_(feed, v);

});

instance() = this;

}

void loop() {

if (!iot::wifiOK()) return;

if (!mqtt_.connected()) {

if (millis() - lastTry_ < 10000) return;

lastTry_ = millis();

user_ = iot::setting("aiouser", AIO_USERNAME);

String key = iot::setting("aiokey", AIO_KEY);

if (user_.isEmpty() || key.isEmpty()) { Serial.println("[AIO] aiouser/aiokey not set"); return; }

String id = "esp32-" + String((uint32_t)ESP.getEfuseMac(), HEX);

Serial.println("[AIO] connecting...");

if (mqtt_.connect(id.c_str(), user_.c_str(), key.c_str())) {

Serial.println("[AIO] connected");

for (int i = 0; i < nSub_; i++) {

mqtt_.subscribe(topic(subs_[i]).c_str());

if (requestLast_) mqtt_.publish((topic(subs_[i]) + "/get").c_str(), "");

}

// subscribe to throttle/error topics so we learn about rate limits

mqtt_.subscribe((user_ + "/throttle").c_str());

} else {

Serial.printf("[AIO] failed, state %d\n", mqtt_.state());

}

}

mqtt_.loop();

}

bool connected() { return mqtt_.connected(); }

// Returns false if not connected or if the per-minute budget is used up.

bool publish(const char* feed, const String& value) {

if (!mqtt_.connected()) return false;

uint32_t now = millis();

if (now - windowStart_ > 60000) { windowStart_ = now; sent_ = 0; }

if (sent_ >= 28) { Serial.println("[AIO] rate budget used, skipping"); return false; }

sent_++;

return mqtt_.publish(topic(feed).c_str(), value.c_str());

}

bool publish(const char* feed, float v, int dp = 2) { return publish(feed, String(v, dp)); }

private:

String topic(const char* feed) { return user_ + "/feeds/" + feed; }

static AdafruitIO*& instance() { static AdafruitIO* p = nullptr; return p; }

WiFiClientSecure tls_;

PubSubClient mqtt_;

FeedCallback cb_ = nullptr;

const char* const* subs_ = nullptr;

int nSub_ = 0;

bool requestLast_ = true;

String user_;

uint32_t lastTry_ = 0, windowStart_ = 0;

int sent_ = 0;

};

}  // namespace cloud